01232 High Temperature Operating Life Chamber (center)

Model Despatch, Type PBC2-18

Watlow controller 988A-12CF-NUDH (with data communications option) serial no. 174305

| Pin | Signal |
| --- | --- |
| 1 | 0-5 VDC+ |
| 2 | 0-5 VDC- |
| 3 | T+/R+ |
| 4 | T-/R- |
| 5 | TRANSMIT |
| 6 | RECEIVE |
| 7 | COM |
| 8 | NOT USED |
| 9 | TC+ |
| 10 | TC- |
| 11 | EARTH GND |
| 12 | DC+ |
| 13 | DC- |
| 14 | COM |
| 15 | mA+ |
| 16 | V+ |
| 17 | mA-, V- |
| 18 | S1, mV-, mA- |
| 19 | S2, V+, T.C.+ |
| 20 | S3, mV+, mA+, V-, T.C.- |
| 21 | L1 |
| 22 | L2 |
| 23 | DC+ |
| 24 | DC- |

The 988A-12CF-NUDH controller is supplied with a "U" board; you can select via the comms menu either EIA-232 or EIA-485 operation. The [IntF] parameter is defaulted to EIA-232. To select the multidrop interface, enter the Setup Menu [SEt]. Use the up-arrow or down-arrow key to advance to the communications Menu [COM]. Press the MODE key until the interface prompt [IntF] appears. Controllers equipped with the EIA-232 interface do not require an interface selection.)

To run a network with multiple devices Watlow uses the ANSI X3.28 Protocol which provides a response to every message. It will also work with the EIA-232 interface, but you are limited to one controller and a host computer.

To select the protocol, go to the Setup Menu [SEt]; use the up-arrow or down-arrow key to advance to the Communications Menu [COM]. Press the MODE key until the protocol prompt [Prot] appears. Select either [FULL], for ANSI X3.28 2.2 - A.3, [On] for XON/XOFF, or [Mod], for Modbus RTU.

If you are using ANSI X3.28 Protocol, choose an address number for each controller using the address prompt [Addr], which follows the protocol prompt [Prot].This prompt will only appear if [Prot] is set to [FULL] or [Mod].

Watlow controller 988A-12CF-NUDH

| 9 8 | single-channel, temperature/process controller, 1 analog input, 1 digital input, 1 analog or digital input, 4 outputs. |
| --- | --- |
| 8 | 100 to 240Vı (ac/dc) nominal vertical mounting |
| A | Standard software |
| 1 | Input 1: Thermocouple only (Excluding Type B, R, and S) |
| 2 | Input 2: Universal signal conditioner |
| C | Output 1: Switched dc or open collector, isolated |
| F | Output 2: Universal process 0-5VÎ (dc), 1-5VÎ (dc), 0-10VÎ (dc), 0-20mA, 4-20mA, isolated |
| N | Output 3: Retransmit, 0-5VÎ (dc), 1-5VÎ (dc), 0-10VÎ (dc) |
| U | Output 4: EIA/TIA-485 / EIA/TIA-232 communications, opto-isolated |
| D | Display (DH not listed) |
| H |  |